## equipt
# | Item | unit | Quantity | Unit Price | Total Price | Specification
1 | Acid dispenser | Pcs | 2 |  | 0 | Chemical Compatibility: The dispenser must be constructed from materials resistant to a wide range of corrosive acids (e.g., concentrated HCl, sulfuric acid, nitric acid). Common materials include PTFE, FEP, PFA, and borosilicate glass components. Volume Range & Accuracy: 0.5-5 mL, 1-10 mL, 5-50 mL. Accuracy and precision are critical, often conforming to ISO 8655 standards, with typical accuracy variations of around ±0.5%. Safety Features: Anti-drip Nozzle Cap: Prevents contact with hazardous reagents and avoids unwanted drops. Recirculation Valve (Priming Valve): Allows air bubbles to be purged from the system and the initial liquid returned to the bottle, minimizing reagent waste and splashes during priming. Secure Bottle Attachment: The dispenser head should fit securely onto the reagent bottle, often including multiple adapters (e.g., 28mm, 32mm, 38mm, 40mm, 45mm) for universal compatibility. Autoclavability: Many lab environments require sterile equipment. The dispenser should be fully autoclavable at 121°C (250°F) and 15 psi to ensure decontamination.
2 | Alcohol meter | Pcs | 6 |  | 0 | Resolution: Alcohol content:0.01vol.% Density:0.00001g/cm3 Specific Gravity:0.00001 Brix:0.02 Babo:0.02 Baume:0.01 Temperature:0.010C Model: D05976 Material: Metal, Rubber Color: Black/optional Size: 198*40mm (7.7"x1.5") Measuring Range: 0~80% ATC: 10~30°C With measuring range 0-80% Easy to focus & calibrate, accurate and reliable With built-in ATC (Automatic Temperature Compensation)
3 | Anaerobic Jar | Pcs | 8 |  | 0 | Anaerobic Jar 3.5 L with White Lid, Clamp, Insert, and Pressure Gauge, pack of 5, Made of polycarbonate. Fitted with Lid having vacuum cum pressure gauge, vacuum/pressure valve and safety valve. The lid of the jar is made of aluminum pressure die casting fitted with pressure valve, Vacuum valve, safety valve & compound pressure cum vacuum gauge, three finger aluminum Clamp and sealing "o" ring and air/gas chuck. - This is recommended only where inside condition of air of jar is removed using vacuum pump and co2 is filled in the jar using co2 gas cylinder to generate anaerobic condition in the jar also supplied with plain aluminum pressure die casted lid which can be fitted when ready made gas packs are used to generate the anaerobic conditions inside the jar
4 | Antibiotic disc dispenser | Pcs | 15 |  | 0 | 6 place, 90mm skirt, Store dispenser in its container with desiccant
5 | Autoclave | Pcs | 11 |  | 0 | Volume: 120L, Power: 4.5kW ,Voltage:AC 220V, 50HZ, Design pressure: 0.25 Mpa Design temperature: 139℃, Rated working pressure: 0.22Mpa, Rated working temperature:134℃, Sterilization temperature: 116℃~134℃ Temp accuracy: ±1℃, Sterilization time setting range: 4~120 min, Chamber volume: Ø450×760 , Bucket dimension: Ø420×660, Basket dimension (optional: Ø410×350×2 , Outer dimension: 555×599×1200
6 | Automatic Hand Dryer | pcs | 4 |  | 0 | Hand Dryer with Accessory Stand & HEPA Filter for, high speed
7 | Automatic mortar grinder accessories fitting sets | pcs | 4 |  | 0 | "Mortar Material Options: Agate / Zirconium Oxide / Tungsten Carbide / Hardened Steel / Stainless Steel / Alumina Mortar Capacity: 250 mL – 500 mL (model dependent), Pestle: Matching material to mortar, wear-resistant, suitable for automatic operation Scraper: Adjustable scraper assembly, PTFE / Stainless Steel / Polyurethane, ensures uniform grinding, Lid / Cover: Transparent, dust-tight, safety-interlock compatible Grinding Pressure System: Adjustable grinding pressure for reproducible results Mounting & Fittings: Complete mounting and fixing accessories, corrosion-resistant, easy installation Compatibility: Fully compatible with specified automatic mortar grinder model Standard Supply: 1 × Mortar 1 × Pestle 1 × Scraper Assembly 1 × Lid/Cover Complete fitting and mounting hardware Application: Grinding, mixing, and homogenization of laboratory samples"
8 | Balance, Electrical | Pcs | 16 |  | 0 | Capacity x Readability: 220 g x 0.01 mg,Repeatability:0-60g ± 0.015 mg / 60-220g ± 0.25 mg,Linearity: ± 0.1 mg, Corner Load (g)* : ± 0.15 mg (100 g),Test Load(g),Minimum Sample Weight (mg): Typically 20 mg minimal initial weighing according to USP,Tare Range: FULL to Capacity,Stabilization Time (avg): < 2 second,Response Time (avg): < 6 second,Sensitivity Drift:(10°-30°C/50°F-86°F: ± 1 ppm / °C,Selectable Weight Units : g, kg, ct, lb, oz, ozt, tlh, tls, tlt, GN, dwt, mg, /lb, tlc, mom, K, tol, bat, MS and parts per pound,Display Update: 0.2-0.4 seconds (Depends on The Filter Level Selected)
9 | Berman’s apparatus | Pcs | 12 |  | 0 | Device Type: Manual parasitological concentration apparatus, Principle of Operation: Sedimentation and migration of helminth larvae and eggs in water by gravity Stand & Support: Metal stand with heavy, stable supporting base. Corrosion-resistant coated or stainless-steel construction Funnel Holding System: Height-adjustable metal fork/clamp. Secure locking mechanism for funnel stability. Funnel: Borosilicate glass with a capacity of 500 mL, and Conical Shape and narrow stem Outlet System: • Fixed glass outlet tube • Adjustable flow-control knob / stopcock (glass or PTFE) Collection System: • Sediment collected into test tube or beaker placed below outlet Operating Mode: Fully manual operation Maintenance & Cleaning: • Easy to dismantle, clean, and reassemble Body / Frame Material: stainless steel stand with glass components Safety & Stability: • Stable base to prevent tipping • Secure clamp to avoid glass breakage Application: Parasitology and veterinary diagnostic laboratories for recovery of nematode larvae and eggs Standard Supply: Berman’s apparatus complete with metal stand, adjustable holding clamp, glass funnel, outlet tube with stopcock, and user manual
10 | Biosafety Cabinet Class-II A2 | Pcs | 6 |  | 0 | Vibration of Work Surface:≤2μm, Sound Emission:≤65dB (A), Illumination:≥900Lux, Standard Accessory:1. Air Velocity Sensor;2. LCD Touch Screen;3. LED Lamp×2;4. UV Lamp×1;5. Explosion-Proof/Waterprof Socket×1;6. Motorized Front Window;7. Footmaster Caster;8. Foot Switch;9. Removable shallow-plate stainless steel work surface;4. Optional Accessory:1. Base Stand;2. Gas tap;3. Water tap;4. Pressure Sensor;6. ULPA filter;7. Upgraded LED Lamp;5. 6. Net Weight:130kg (Without base stand 9. ), 140kg (With base stand), 7. Gross Weight:165kg (Without base stand), 175kg (With base stand),8.Shipping Size (W×D×H)(mm):850×990×1640. 9. Shipping Volume: 1.4 m³
11 | Biosafety Cabinet Class-III | Pcs | 3 |  | 0 | "Type: Class III, HEPA Filter: ULPA Filter 2, ULPA Filter: Air Supply Filter: Efficiency 99.999% at 0.12 µm, First exhaust filter: Efficiency 99.999% at 0.12 µm, Second Exhaust Filter: Efficiency 99.999% at 0.12 µm Airflow Volume 440-870 cfm Pressure: ≥ negative 120 Pa Pre-Filter: Polyester fiber, Washable LED Lamp: 28Wx2 UV Lamp: 30Wx3, 8Wx1 Illumination: ≥1000 lux Inner Dimension (WxDxH): 1240x650x650 mm Overall Dimension (WxDxH): 1765x855x2100 mm Pass Box Size (WxDxH): Internal Size: 400x390x340 mm External Size: 480x400x400 mm Package Size: 1850x1080x2280 mm Front Window: 8 mm Toughened glass, Anti-UV Chamber Material: Epoxy coated cold-rolled steel Noise Level: NSF 49≤61 dB / EN 12469≤48 dB Weight: 395 kg Power: 700 W Power Supply: 110V~240V/50Hz~60Hz Accessory: Fluorescent Lamp; Remote Control; Drain Valve; Pressure Meter; Gloves; UV Lampx4 "
12 | Bunsen burner | Pcs | 12 |  | 0 | Portable Laboratory Bunsen Burner
13 | Caliper (skin reader for PPD test) | Pcs | 8 |  | 0 | Mechanical and electronics TB Caliper, 0 to 6" Range, 0.01mm Resolution
14 | Centrifuge, Heamatocrit | Pcs | 13 |  | 0 | Centrifuge, Micro-Heamatocrit, direct reading, 24 tube capacity: Digital Hematocrit Centrifuges - Digital speed and time display. - 10 memory programs for presetting parameters. - Auto accelerating and decelerating functions. - Max. Speed (rpm) 1,000~12,000 - Max. RCF (g) 14,810 - Max. Capacity (ml) Hematocrit x 24 - Timer (min:sec) 99:59 - Weight (kg) 10 - W x D x H Size (cm) 29 x 32 x 25 Accessories: - 2 Rotor (Capacity: Hematocrit x 24pcs; Max. Speed (rpm): Max. 12,000rpm, Tube size (Ø x L) mm: 1.1~1.2 x 75) - 10 pack HT capillary tubes
15 | CO2 gas regulator | pcs | 3 |  | 0 | Gas Type: Carbon Dioxide (CO2CO sub 2 CO2), Inlet Pressure (Max):.Up to 300 Bar (approx. 4000 PSI), Outlet Pressure (Max):.0-4 bar (welding/industrial) to 0-10 bar (aquarium/specialty) or 0-90 PSI (mini regulators),Flow Rate: 0-25 L/min (welding) up to 96 m3/hm cubed / h m3/h (industrial) Stage Type Single-stage:.Single or Dual (Two-stage)
16 | Colony counter | Pcs | 10 |  | 0 | Height adjustable:illuminated area of 145 mm Ø with direct or,indirect glare free lighting, frosted glass and clear,glass plate with cm 2 und 1/9 cm 2 graduation, electrical contact pin with felt tip pen for marking,Petri dishes of up to 145 mm Ø can be used.
17 | Deionizer | pcs | 3 |  | 0 | Treated Water Specifications-Flowrate : 2.0 l/min max, Product water capacity: 12 liter per hour,-Inorganics: 18.2 MΩ-cm ,-TOC: 1– 3 ppb Bacteria : <0.1 CFU/ml,-Bacterial endotoxin : <0.001 EU/ml, -pH : Effectively neutral ,-Particles : Ultrafiltration , -RNase and DNase - - Removed -Cartridge capacity (LC162) 45,000 liters >,18MΩ-cm per single purification pack/μS at pH 7.0, 70,000 liters >1MΩ-cm per single pur, Silica 2 ppm max.Temperature 1 - 40°C , Electrical Requirements , Mains input 100 - 240V ac, 50 - 60Hz , Reservoir level connection: Jack Plug 3.5mm Noise level during recirculation: <40dBA
18 | Diamond pen | Pcs | 7 |  | 0 | 6" holder with artificial diamond (hard stone) embedded at one end with screw cap to mark on microscope glass slides
19 | Duran bottle 250ml | Pack of 10 | 7 |  | 0 | Capacity: 250 mL nominal volume,,Material: Borosilicate glass, high chemical and thermal resistance Neck & Closure: Standardized threaded neck and compatible with screw caps, stoppers, or seals Body Design: Cylindrical shape for stability and ease of handling, Smooth surface for easy cleaning and Graduated markings Temperature Resistance: Thermal shock resistant: up to ~150 °C sudden temperature change and autoclavable: up to 121 °C (15 min, standard autoclave conditions) Chemical Resistance: Resistant to acids, bases, and most organic solvents Operating Conditions: Room temperature to autoclave conditions
20 | Duran bottle 500ml | Pack of 10 | 7 |  | 0 | Capacity: 500 mL nominal volume,,Material: Borosilicate glass, high chemical and thermal resistance , Neck & Closure: Standardized threaded neck and compatible with screw caps, stoppers, or seals Body Design: Cylindrical shape for stability and ease of handling, Smooth surface for easy cleaning and Graduated markings Temperature Resistance: Thermal shock resistant: up to ~150 °C sudden temperature change and autoclavable: up to 121 °C (15 min, standard autoclave conditions) Chemical Resistance: Resistant to acids, bases, and most organic solvents Operating Conditions: Room temperature to autoclave conditions
21 | Electro conductivity meter | Pcs | 3 |  | 0 | Device Type: Portable / Benchtop Electroconductivity Meter Measurement Range: • EC Range: Typically 0–20,000 µS/cm (or as specified) • TDS Range: 0–10,000 ppm (if applicable) • Temperature Range: 0–100 °C (if equipped with temp compensation) Probe / Sensor: • Type: Conductivity cell probe (fixed or replaceable) • Cell Constant: 1.0 / 1.5 / 2.0 cm⁻¹ (as specified) • Material: Stainless steel / Titanium electrodes Display: Digital LCD / Backlit screen (as specified) Calibration: • Method: Manual / Automatic calibration • Points: 1‑point / 2‑point (as specified) • Standards: Calibrate with standard conductivity solutions Temperature Compensation: • Type: Automatic / Manual • Reference Temp: 25 °C (usually) Accuracy: ±1% of reading (typical, as specified) Resolution: • Conductivity: 0.1 µS/cm • TDS: 0.1 ppm • Temperature: 0.1 °C (as specified) Power Supply: Battery powered / Rechargeable / Adapter (as specified) Output & Interfaces: USB / RS232 / Bluetooth (if equipped, as specified) Body Material: ABS Plastic / Rubberized grip / Water‑resistant housing (as specified) Operating Conditions: 0–50 °C, <95% RH (typical, as specified) Safety & Protection: • Over‑range indication • Probe protection cap • Auto power‑off (if available) Application: Water quality testing, environmental monitoring, laboratories, aquaculture, hydroponics Standard Supply: EC meter with probe, batteries, calibration solutions, manual, carrying case
22 | ELISA Reader(Elisa reader filter for multi scan) | pcs | 7 |  | 0 | """Linear Measurement Range : 0.000~3.500 Abs Absorbance Units(A) Plate Type: 96-well plate Light Source: Monochromatic LED Reading Speed: 5s with single wavelength for whole plate Input: Mouse, membrane keyboard, touch screen Working conditions: Humidity: 15%~85% Temperature: 10°C~35°C Measure System: 8~channel optical system Band Width and Accuracy: 8 nm ± 2 nm Shaking: 3 shaking modes with variable time Wavelength: 405, 450, 492, 630nm, two more optional LED light source with lifespan over 10,000 hours. Repeatability:<1% Stability: <0.005 Abs/10 min Linearity: R≥0.995 Display: 7 inch TFT LCD, 800x480 Pixels Interface: RS~232, 4 USB ports, SD Card Printer: Built-in printer, external printer Power Supply: AC 100~240V, 50/60 Hz Dimension: 400Lx310Wx169H mm Weight: 6 kg"""
23 | Graduated beaker 150ml | Pack of 10 | 4 |  | 0 | Borosilicate glass, graduated 150 ml
24 | Graduated beaker 500ml | Pack of 10 | 4 |  | 0 | Borosilicate glass, graduated 500ml
25 | Graduated cylinder, 100ml(plastic) | Pack of 12 | 4 |  | 0 | Plastic, graduated, 100ml
26 | Graduated cylinder, 250ml(plastic) | Pack of 4 | 4 |  | 0 | Plastic, graduated,250ml
27 | Heat block | pcs | 3 |  | 0 | Controller Type:- Feedback Control PID, Capacity (Metric):- 1.5 mL ,Lighting: LED, No. of Places:- 96, Plug Type:- Cord set w/various plugs, Temperature Range (Metric): Ambient +5°C to 100°C, Type: Platform Shaker, Display: Digital, Speed:- 150 to 1,500rpm (96-well block); 150 to 1,200rpm (0.5mL, 1.5mL tube blocks), Temperature Accuracy:- 0.5°C (0.9°F) at 80°C, Warranty:- 2 Years, Dimensions (L x W x H):- 13.1 x 10.9 x 6.7 in. (334 x 276 x 170 mm), Diameter (Metric) Orbit:- 2 mm, Shipping Weight (English):- 18.3 lb.,Shipping Weight (Metric):- 8.3 kg ,Voltage:-230 V ,Frequency:-50/60 Hz ,Unit Size, Type, platform shaker, No. of Places, 24
28 | Heating mantle for flasks 1-5 litter | Pcs | 1 |  | 0 | Temperature range:up to 450 deg C, Speed range up to 1200 RPM,For vessel capacity 5Liter, Power 700 watt, Separate control switch for heat and stirring
29 | Hematocrit reader (PCV) | Pcs | 8 |  | 0 | Flow Control: Designed for smooth and consistent ink delivery
30 | Hematology analyzer | Piece | 4 |  | 0 | """Veterinary auto hematology analyzer WBCi/HiGB • Measuring Principles: Impedance method for cells counting Cyanide free colorimetric method for HGB measurement • Measuring Channels : Double channel for counting • Throughput: From beginning of a test to completion of the printout: :≤60 Seconds • Counting Modes: Whole Blood, Pre-diluted Blood • Sample Volume: Whole Blood: ≤ 14µL Pre-diluted Blood: :≤20µL • Species of Animals: 13 pre-determined animal settings (cat, dog, horse, mouse, rat, rabbit, pig, cow, buffalo, monkey, camel, sheep and goat ) plus 3 programmable. • Parameters & Histograms: Three-part differentiation of WBC covering dog, cat, mouse, rat, rabbit, monkey and horse, 21 parameters + 3 histograms WBC, Lymph#, Mid#, Gran#, Lymph%, Mid%, Gran%, RBC, HGB, MCV, MCH, MCHC, RDW-CV, RDW-SD, HCT, PLT, MPV, PDW, PCT, P-LCR, P-LCC, WBC Histogram, RBC Histogram, PLT Histogram • Alarm System: Multiple alarms of running status including blank counting abnormal, reagents empty, waste full, counting hole clogged, counting area bubble, recorder abnormal, external printer abnormal and circuit abnormal, etc • Measures of Unclogging Treatment: Back flushing, soaking and burning. Automatic unclogging treatment before and after the sample analysis • Printout: English print report, optional report formats • Storage: Can store more than 10,000 sample results, including parameters, histograms and animal information. Support multiple conditions combination query • Communication: Open communication protocol. Support connection to LIS • Data Import & Export: Import or export data via USB port • Quality Control System: Include 9 QC files, automatically draw quality control chart, provide the average, SD, CV • Calibration System: Support manual and automatic calibration • Software Upgrading: Upgrading software via USB port • Display & Touch Screen: 10.4 inch true color TFT LCD and touch screen • Printer: Built-in patented structure thermal printer, paper 80mm wide, easy to load paper. Report printing takes only 5 seconds • Data Ports: One network port, one RS-232 port, four USB ports • Optional External Devices with USB Port: Standard 101 keyboard, mouse, barcode scanner and multiple printers • Power: 220V, 50Hz Input power: ≤150VA • Operating Environment: Environment Temperature: 10°C ~ 30°C Relative Humidity: ≤ 70% Atmospheric Pressure: 70.0kPa ~ 106.0kPa • Dimensions: 368mm x 448mm x 4 75mm • Weight: ≤23Kg Reagents: Provide original start up reagents"""
31 | Hemostatic Forceps Curved | Piece | 14 |  | 0 | Safety & Durability Features: Impact-resistant body, secure crystal top, leak-proof ink cartridge
32 | Hemostatic Forceps Straight | Pcs | 12 |  | 0 | Stainless steel, 5.5 inch,
33 | Hydrometer(humidity, T0 & time) | pcs | 3 |  | 0 | Hygro-Thermometer with Humidity Alert • Large, easy-to-read triple LCD displays % Relative Humidity, Temperature and Dew Point • Measurement ranges: - Humidity: 1 to 99%RH - Temperature: 14 to 122°F (-10 to 50°C) • Basic Accuracy of 4%RH (35% to 75%), ±5%RH (<35% or>75%), ±1.8°F/1°C • Max/Min with Reset • Beeper sounds and bright LED flashes when humidity level exceeds user preset (50, 60, or 65 %RH) • Switchable °F or °C unit • Desktop stand can be removed for wall mount use • Dimensions: 3.7 x 3.7 x 0.9” (94 x 94 x 23mm); Weight: 4.8oz (136g) • Complete with stand and 2 AA batteries"
34 | Microplate washer, Automatic | Pcs | 10 |  | 0 | Washing Head 8 needles /12 needles,Washing Channel Standard 2 channels, can be selected according to requirements, up to 4 channels,Washing Mode Plate or strip, Washing Times 1-99 times adjustable, Washing Strips 1-12 strips adjustable, Dipping Time 0-24 hour adjustable, Shaking Time 0-24 hour adjustable, Residual Volume ≤1μL/well,Liquid Volume 10-3000μL, adjustable, step by 1μL, Sipping Time 0.1-9.9s, Washing Method With 2-point liquid suction, bottom washing, and fine adjustment of wash liquid pressure,External Dimension (L*W*H) 446*368*164mm, N.W./G.W. 9kg/14kg, Shipping Dimension (L*W*H) 600*490*530mm, MRW-H330 or similar
35 | Co2 Incubator | Pcs | 7 |  | 0 | Chamber Volume（L） 160, Temperature Range（℃） RT+3～60, Temperature Stability（℃） ≤±0.2, Temperature Uniformity（℃） ≤±0.3, Timing Range 1～9999mins or without timing, CO2 Range 0～20%, Humidity Method Natural Vaporization, Power Supply AC220V, 50Hz, Power Rating (W） 900, Chamber Size (W×D×H)cm 50×50×65, Exterior Size(W×D×H)cm 69×69×103, Package Size (W×D×H)cm 70×78×118, Net/Gross Weight（kg 75/110, CelMate® CO₂ Incubator- CO₂ Control, UV Lamp, Direct heat and air jacket, Infrared (IR) CO2 Sensor,ULPA Filter, ISO Class 5 work zone,Validated 90°C moist heat decontamination cycle, UV Lamp,FDA-listed, Class II, 510k exempt medical device, Available volume: 170 L
36 | Incubator (Microbiological) | Pcs | 6 |  | 0 | Made of: Textured stainless steel casing and stainless steel interior; Housing: Rear zinc-plated steel; Interior: Made of stainless steel, easy to clean interior, reinforced by deep drawn ribbing with integrated and protected large-area heating on four sides; Door (outer): Fully insulated stainless steel doors with 2-point locking (compression door lock); Door (Inner): inner glass doors; Installation: on lockable casters; Internals (Shlefs): 2 stainless steel grids and shoul have additional 2 grids for possible adjustment; Volume: Not less than 450 Litter; Control System: Control COCKPIT: adaptive multifunctional digital PID-microprocessor with 2 high-definition TFT-color display; Temprature: 2Pt100 sencors Class A in 4 wire circuit, mutually monitoring and taking over the performance at the same temprature value; Timer: Digital backwards counter with target time setting, adjustable from 1 minute to 99 days; Temprature: minimum 10 oC above ambient up to +80 oC; Resolution of dispaly for set point and actual temprature values: 0.1 oC; Control technology: Adjustable parameters: temprature (celsius or Fahrenheit), fan speed, air flap position, program time, time zones, daylight savings time; Function Heat Balance: adapting the distribution of the heating perfomrance of the upper and lower heating ciruit from -50% to +50%; Function Setpoint Wait: the process time does not strat until the set temprature is reached; Should have Language Setting; Calibration Points: at least three freely selectable temprature values; Ventilation System: Fan: forced air circulation by quite air turbine, adjustable in 10% steps for each segment individually; Fresh air admixture: adjustment of pre-heated fresh air admixture by air flap control in 10% steps for each segment individually; Vent: Vent connection with restrictor flap; Communication: Programming: AtmoCONTROL software on a USB stick for programming, Managing and transferring programs via Ethernet interface of USB port; Documentation: Program stored in case of power failure; Safety: Alarm: audible and visual; Temprature Control: Overtimeprature monitor TWW, Protection class 3.1 or adjustable temprature limiter TWB, Protection class2, selectable on display; AutoSAFETY: additionally integrated over and under temparture monitor "ASF", automatically follwoing the set point value at a preset tolerance range, alarm in case of over or undertemprature, heating is switched off in case of over temprature; Temprature Control: mechanical temprature limiter TB, Protection class 1 according to DIN 12880 to switch off the heating approximatly 20 oC above nominal temprature; Autodiagnostic System: for fault analysis; Power Supply: 220-230 V, 50/60 Hz, minimum 2meter power cord with two round ended plug (Type F); Should be delivered with: Certeficate of compliance, Calibration certificate at least for +37 oC or additional point, English version User Manual (bothe Hard and soft copy);
37 | Incubator (Microplate shaker) | Pcs | 7 |  | 0 | Temp. setting range 0℃~80℃, Temp. control range RT+5℃~80℃, Time range 1min ~ 99h59min/∞, Temp. accuracy ≤±0.5℃, Display accuracy 0.1℃ Temp. uniformity ≤±0.5℃,Speed range 200~1350rpm, Orbit 3mm Heating time ≤10min(from 25℃ to 80℃), Auto preheating Yes, Auto resume to run Yes Capacity 6 Microplates or Deepwell plates, Power consumption 300W,Electricity 220V or 110V AC /50-60HZ Fuse 250V 3A Ф5×20 External dimension(mm)（W*D*H） 320x340x200, Net weight 8-11kg
38 | Inoculation wire loop | Pack | 30 |  | 0 | Nichrome/metal wire loops on handles, calibrated, 10µL, Twisted Nichrome V, 2 mm with Insulated Handle, pack of 5
39 | Laptop Computer | Pcs. | 12 |  | 0 | 1. Chipset & Performance: Processor: Apple M2 chip (2nd Ge, 8-core CPU (4 performance + 4 efficiency co, 10-core GPU (base) or 12-core GPU (upgrade) 16-core Neural Engine, Memory: 24GB installed memor 100GB/s memory bandwidth• Storage: 1TB SSD , SSD speeds up to 3.5GB/s 2. Display: Size: 13.3-inch Retina (diagonal, Resolution: 2560 × 1600 (227 PPI), Brightness: 500 nits, Color Support: P3 wide color, True Tone, Refresh Rate: 60Hz (fixed, no ProMotion). 3. Design & Ports, Dimensions: 11.97 × 8.36 × 0.61 inches, Weight: 3.0 lb (1.4 kg), Ports: 2 × Thunderbolt 3 / USB-C (supports DisplayPort), 3.5mm headphone jac, No HDMI, SD card slot, or Magf Keyboard: Magic Keyboard with Touch ID (no Touch Bar) 4. Battery & Power: Battery Life: Up to 20 hours video playback, Adapter: 67W USB-C power adapterFast Charging: Not officially supported (no higher-wattage option). 5. Audio & Camera: Speakers: Stereo (no high-end six-speaker system, Microphones: Dual-mic array (studio-quality on higher mode l, Webcam: 720p FaceTime HD (not 1080p). 6. Wireless & OS, Wi-Fi 6 (802.11ax) + Bluetooth 5., macOS Sonoma pre-installe , Security: Touch ID, Secure Enclave
40 | Loop holders(handel) | Pcs. | 16 |  | 0 | Insulated Brass Holder, 8 inch For loops with out handle
41 | Macmaster egg counting chamber | Pcs. | 12 |  | 0 | For veterinary medicine (e.g. Helminth eggs counting) with 2 counting areas, cover glass with 2 counting grids glued on 3 supports. With counting area of (10 x 10 x 1.5 mm) and dimension of approx. 76 x 32 mm.
42 | Magnetic hot plate stirrer | pcs | 9 |  | 0 | Magnetic Stirrer Hot Plate Stirrer LCD Display 300℃ Max.5L 5.9inch Magnetic Stir Plate 1600RPM Magnetic Mixer Magnetic Hotplate Stirrer with Stir Bar&Temp Probe Senso,220-230V
43 | McFarland densitometer | Pcs | 6 |  | 0 | To measure turbidity in the range of 0.0-6.0 units (0-180 x 107 cells/ml). Light source-LED, Accuracy 99%, Measurement time-1s, sample volume- <2ml, dimension (wxlxh) 165x115x75mm, display type-LCD, weight-0.7kg, power consumption-12v,7MA/0.1W, power supply 220-240voltage.
44 | Micropipette (Multichannel , 12 channels, 5 to 50 µl) | Pcs | 19 |  | 0 | Precision pipette suitable for volume 5-50 µL with pipette stand
45 | Micropipette (Multichannel , 12 channels, 50 to 100 µl) | Pcs | 7 |  | 0 | Precision pipette suitable for volume 50-100 µL with pipette stand
46 | Micropipette (Multichannel , 12 channels, 50 to 300 µl) | Pcs | 12 |  | 0 | Precision pipette suitable for volume 50-300 µL with pipette stand
47 | micropipette (Single channel , 0.5-10 µl) | Pcs | 16 |  | 0 | Precision pipette suitable for volume 0.5-10 µl with pipette stand
48 | micropipette (Single channel , 10-100 µl) | Pcs | 16 |  | 0 | Precision pipette suitable for volume 10-100 µl with pipette stand
49 | micropipette (single channel, 100-1000 µl) | Pcs | 21 |  | 0 | Precision pipette suitable for volume 100-1000 µl with pipette stand
50 | micropipette (Single channel, 50-200 µl) | Pcs | 12 |  | 0 | Precision pipette suitable for volume 50-200 µl with pipette stand
51 | micropipette (Single channel, 5-50 µl) | Pcs | 7 |  | 0 | Precision pipette suitable for volume 5-50 µl with pipette stand
52 | Micropippette (Multichannel, 100-300 µl) | pcs | 8 |  | 0
53 | micropippette (single channel, 25 - 1000 µl) | pcs | 7 |  | 0
54 | Microplate reader | Pcs | 7 |  | 0
55 | Microplate washer, Automatic | pcs | 6 |  | 0 | Washing Head 8 needles /12 needles,Washing Channel Standard 2 channels, can be selected according to requirements, up to 4 channels,Washing Mode Plate or strip, Washing Times 1-99 times adjustable, Washing Strips 1-12 strips adjustable, Dipping Time 0-24 hour adjustable, Shaking Time 0-24 hour adjustable, Residual Volume ≤1μL/well,Liquid Volume 10-3000μL, adjustable, step by 1μL, Sipping Time 0.1-9.9s, Washing Method With 2-point liquid suction, bottom washing, and fine adjustment of wash liquid pressure,External Dimension (L*W*H) 446*368*164mm, N.W./G.W. 9kg/14kg, Shipping Dimension (L*W*H) 600*490*530mm, MRW-H330 or similar
56 | Microscope, Camera-Fitted | Pcs | 15 |  | 0 | "Condenser numerical aperture: NA = 1.25; Binocular Digital Microscope with Seidentopf Binocular Head, Wide Field Eyepiece, Backward Quadruple Nosepiece, Achromatic Objectives, LED Illumination and 3.2MP Camera and Software Viewing Head: Seidentopf Binocular Head, Inclined at 30°, Interpupilary Distance 48-75mm Camera System: Resolution:2048×1536 (3.2Mega Pixel) Output Mode: USB2.0 Operation System: Window 2000/ XP/ Vista/ Win 7/ Win8/ Mac Software: ScopeImage 9.0 Frame Rate: 9 fps @ 2048×1536 Eyepiece: WF10×/ 20mm Nosepiece: Backward Quadruple Nosepiece Objective: Achromatic Objective 4×, 10×, 40×, 100× Stage: Double Layers Mechanical Stage 132×142mm/ 75×40mm Condenser: Abbe NA1.25 with Iris Diaphragm & Fliter Focusing: Coaxial Coarse & Fine Adjustment, Fine Division 0.002mm, Coarse Stroke 37.7mm per Rotation, Fine Stroke 0.2mm per Rotation, Moving Range 20mm Illumination: 3W LED illumination, Brightness Adjustable, protective cover "
57 | Microscope, Stereomicroscope | Pcs | 12 |  | 0 | Binocular Stereo Microscope; WF10×/Φ22mm Eyepieces; 6.7×-45× Magnification, Zoom objective: 0.67×-4.5×; • Stereo microscope for home or classroom use enables students and hobbyists to inspect biological specimens, rocks, stamps, jewelry, and large specimens that require handing or manipulation Optical system: Greenough Optical System Viewing Head: Binocular head, inclined at 45°, rotatable 360°, Interpupillary Distance 54-75mm, diopter adjustment on both eyepieces Zoom Ratio: 1:6.7 Working Distance: 105mm Focusing: Coarse focus knob with tension adjustable, moving range: 106mm. Focusing arm diameter: 76mm Column Height: 300mm Microscope Base Size: 205×275×40mm Illumination: Reflected illumination: 5V 3W LED Transmitted illumination: 5V 3W LED Package 1pc/carton, 51cm*37cm*31cm, Net weight: 5.5kg, Gross weight: 6.5kg
58 | Mortar & Pestle(large) | Pcs | 12 |  | 0 | 1000 ml capacity with 216 mm diameter, White porcelain, pack of 6
59 | Mortar & Pestle(small) | Pcs | 22 |  | 0 | 130 ml capacity and is 100 mm diameter, White porcelain, pack of 6
60 | OmniLog Combo Plus ID System including database ( for identification of veterinary microorganism ) | Pcs | 2 |  | 0 | All Biolog Automated Systems include the OmniLog® GEN III Data Collection Software, Retrospect 2.0, Data Management Software, electronic user guide, OmniLog reader/incubator, computer, printer,pipettor, turbidimeter and a one year warranty.OmniLog Phenotype MicroArray™ (PM) Software.GEN III OmniLog Combo Plus System, UK style plug (220-240 VAC) , GEN III Identification Database: 1. GEN III Database (for both Gram-negative & Gram-positive bacteria), 2. AN Database Requires MicroStation or OmniLog Plus System, 3. YT Database:Requires MicroStation or OmniLog Plus System, 4. FF Database Requires MicroStation or OmniLog Plus System, Turbidity Standard (65%T) for GEN III, 85%T, AN:65%T, (YT: 47%T),FF: 75 %T),Identification Consumables: GEN III microplate, AN mcroplate, FF microplat, YT Microplate, Inoculating fliud-A, IF-B, IF-C1, IF-C2, AN filude, 1250ul Pipette Tips – sterile, racked, box of 960, Reservoirs – sterile, case of 100, Biolog Universal Growth Agar (500g jar, Installation, training and one year waranty.
61 | Oven (Hot air oven, dual purpose) | Pcs | 4 |  | 0 | Capacity: 136 L, Temperature Range a. As incubator : RT+5 °C - 80 °C b. As oven : 80 °C - 220 °C,Shelves 2 pcs Display Resolution 0.1 °C Temperature Accuracy a. ± 0.5 °C b. ± 1 °C Ambient Temperature 5 °C - 35 °C Timing Range 1 - 9999 min Power Consumption 2050 W Power 220 V / 50 Hz Interior Dimension (W x D x H) 550 x 450 x 550 mm Exterior Dimension (W x D x H) 840 x 570 x 730 mm355 L; Standard temperature range from ambient +10°C to 100°C; Fluctuation typically: +/-.25°C; Calibrated PID controller; Overheat safety thermostat; Fan circulation; CE Certified. Microprocessor temperature air controller that ensures precision and reliability Forced air-convection Double layer glass door Large viewing window Designed with polished stainless steel chamber Space between the shelves in the chamber can be adjusted Independent temperature-limiting alarm Semicircular arcs at corners for easy cleaning"
62 | Petridish(large) | Pack of 100 | 20 |  | 0 | Boro silicate (Autoclavable), base diameter 54mm, cover diameter 55mm, size 94/16 mm pack of 10
63 | Petridish(small) | Pack of 100 | 18 |  | 0 | Boro silicate (Autoclavable), base diameter 54mm, cover diameter 55mm, size 60X20mm pack of 10
64 | PH meter (with conductivity Measurement) | Piece | 4 |  | 0 | Meter style : Portable,Dimensions( W)x (H) x (L) (7-1/4) x (3-1/2) x (9),Range (pH): -2.00 to 19.99, Range (mV): ±1999, Range (Temperature °C): -5.0 to 105,Resolution (pH): 0.01,Resolution (mV): 0.1, 1,Accuracy (pH): ±0.01,Accuracy (Temperature): ±0.1, ±1 digit °C,Accuracy (mV): ±0.5, 0.1,Temperature compensation : Automatic or manual,Display : LCD, Interface : RS-232,Waterproof,Power: 220V-240 V, 50/60 HZ, Plug Configurations Type F,Datalogging (points): 500,Certification/Compliance,
65 | Pyrex culture tubes with screw cap | Pack of 100 | 5 |  | 0 | Pyrex, screw neck, inert PTFE liner and rubber disc resistant to repeated dry heat and steam sterilisation cycles (121°C for 2 hours)Ideal for culture work, 20 x 100mm, cap size 20mm, wall thickness
66 | Pyrex culture tubes with screw cap | Pack of 100 | 100 |  | 0 | PYREX 11 mL Screw Cap Culture Tubes with PTFE Lined Phenolic Caps, 16x100 mm, Bottom Style:Round, borsilicate glassTop Style:Threaded, Reusable, 144per pack,
67 | Refrigerator, Digital -80c0 | pcs | 4 |  | 0 | Capacity 816 liter, Cu. ft. :28.8፣Cry boxes (2 in.): 600 Electrical: 230V/50Hz) Amps/breaker (plug type) : 6.1/10 (European) Max shelf weight 245kg, lbs. 111 Interior dimensions H x W x D in. (cm) 51.2 x 34.4 x 28.3 (130.1 x 87.3 x 71.9) Exterior dimensions H x W* x D in. (cm) 78.0 x 39.6 x 38.5** (198.1 x 100.6 x 97.7) Shipping weight lbs. (kg) 712 (323) 820 (372)
68 | Refrigerator, Digital -20c0 | Piece | 7 |  | 0 | The refrigerator/freezer has an independent door for both the refrigerator section and the freezer section. Overall capacity is 18.2 cubic feet. Refrigerator temperature ranges from 1° to 12°C (34° to 54°F); freezer temperature ranges from -10° to -21°C (14° to -6°F). With ABS plastic interior, and textured powder coated steel exterior. Automatic defrost only. Two-Door laboratory Refrigerator/Freezer; Chamber dimension refrigerator section : 18”(W) X 26”(D) X 38” (H) or (457 x 660 x 965) mm; Chamber dimension freezer section15”(W) X 23”(D) X 17.25” (H) or (381 x 584 x 438 ) mm; Total volume cu. Ft : 18.2 ; Exterior dimensions 29.6”(W) X 29.88”(D) X 66.13” (H) or (752 x 759 x 1680) mm
69 | Round bottom flask 500ml, borosilicate glass | Pcs | 10 |  | 0
70 | Scalpel handle | Pcs | 10 |  | 0 | Scalpel Handle # 4, Proof Stainless Steel Scalpel Knife Handle, Lightweight and Durable, Fits Surgical Blades No. 20, 21, 22, 23, 24 and 25
71 | Slide dispenser | Pcs | 10 |  | 0 | Made of ABS plastic, with capacity 80 standard slides 76x26mm, dimension 120x100x165mm. Dispense slides one by one from dust free storage
72 | Slide storing box(Small) | Pcs | 3 |  | 0 | Storage box for 450 slides with a dimension of 34x23x10 cm. Support for 75 slides
73 | Spoon with Spatula | Pcs | 10 |  | 0 | Length: 10.6 inch, Width: 2.5 inch, Pack of 100
74 | Staining jar | pcs | 4 |  | 0 | Convenient kit for slide staining, storage and transportation, Dimensions - 100 x 85 x 55mm deep, a push-fit lid and small adapter to aid drainage; Has a folding handle for transfer and immersion of slides, Slotted and numbered for 25 slides; Fits snugly into the trough whilst the folding handle permits closure of the trough lid during
75 | Steam sterilizer | pcs | 4 |  | 0
76 | Test tube rack | pcs | 8 |  | 0 | Stackable, alphanumeric markings, Autoclavable at +121 °c/+249.8 °f, suitable for water baths. With a dimension of (265 x 126 x 75mm) and designed for 90 test tubes
77 | Automatic tissue processor | pcs | 4 |  | 0 | • Easy loading and full process control by rotating table principle,• Integrated UPS system protects your samples,• Up to 240 cassettes per run (optional two baskets) • Intuitive control panel • System status monitor • Password protection • Delayed start function • Fume extraction system with hood (optional) • Vacuum function (optional) • Remote Alarm System (optional) Specification • Nominal supply voltage / frequency: 100 – 240 V AC (+/- 10%) / 50 – 60 Hz • Power draw: 250 VA • Protective class: (1) I • Power fuses: 2x T 6,3 A • Pollution degree : (1) 2 • Overvoltage installation category: II • Maximum heat emission: 250 J/s • Operating temperature range: +10 to +35°C • Temperature range during storage: +5°C to +55°C • Relative humidity: max. 80% non-condensing • Humidity during storage: < 8% , According to IEC 1010, EN 61010 • Dimensions (W x D x H): 850 mm x 850 mm x 570 mm • Reagent beakers: o Standard: 2000 ml, white, polyacetal o (Optional): 2000 ml, DURAN ® glass o (Optional): 2000 ml, aluminum • Heated Paraffin beakers: o Material : aluminum, double walled • Transport basket: o Standard: 1x Stainless steel, capacity 120 cassettes o Optional: 2x stainless steel, capacity 120 cassettes each • Programming: 20 programs (freely programmable), password protected • Incubation time (Exposure): 1 sec up to 99 h 59 min 59 sec (freely programmable) • Drip of time: 30 sec up to 59 min 59 sec (freely programmable) • Paraffin beaker temp.: Temperature range 40 °C to 80 °C (freely programmable) • Instrument operation modes: Immediate start o Fixed start time: delay programmable by to 999 hours o Fixed end time: process end time delayed by up to 99 hours o Quick start (factory programmed protocols) • Agitation: Yes (factory setting for intensity) • Power failure failsafe position: freely programmable for any protocol step • Control elements: LCD display System Status Monitor • Vacuum function (optional): Pressure difference 0,5 bar maximum (500 hPa) Tissue Processor MTP Includes o Basic unit o Integrated UPS system o 10x reagent beakers (POM, white), 2000 ml o 2x electrically heated paraffin beakers, 2000 ml o 1x transport basket (capacity: 120 cassettes), o stainless steel o Operation manual o The device will include its own dedicated UPS or Stabilizer Basic component / Accessories included • Fume hood for MTP/MSM o Adjustable, transparent fume hood (perspex) o incl- 5 distance bolts • Exhaust fan system incl. carbon filter for MTP/MSM o integrated fan for fume extraction in MTP/MSM o carbon filter system • Extension TWIN BASKET for MTP o Modification of MTP to 9 POM beakers and 3 heated paraffin beakers o Modification of MTP to 2 transport basket holder o Modification of MTP to 2 transport baskets (stainless steel) o Increased capacity of up to 240 cassettes per run
78 | Vortex mixer | pcs | 9 |  | 0 | Equipped with Digital Speed display, Continuous and Touch modes Speed range: 200 to 3000rpm, Speed Accuracy: ±5% rpm, Convenient Speed adjustments by knob control, Stable for long-term operations, Multiple accessories for different vessel
79 | Washing bottle 500ml | pcs | 10 |  | 0 | Plastic, 500ml, capped
80 | Water bath (digital) | pcs | 4 |  | 0 | Classification Specifications 2 rows & 4 holes, Heating Mode Nature water convection heat transfer Function Temp. Range RT+5～100℃ Temp. Resolution Ratio 0.1℃ Temp. Motion ±0.5℃, Temp. Uniformity ±1.0℃ Structure Inner Chamber Stainless steel Outer Shell Cold-rolled steel plate, Surface electrostatic spraying Heater Stainless steel heating tube Power Rating 1.0kW Controller Temp. Control Mode PID Temp. Setting Mode Touch button setting Temp. Display Measuring temperature: displayed on the upper row of the three-digit LED; Mode Setting temperature: displayed on the lower row of the three-digit LED Timer 0～9999min（with timing wait function） Operation Function Fixed temperature operation, timing function, auto stop. Program Mode Optional Additional Function Deviation Correction, menu key locked, power failure back-up, power-off memory Sensor NTC Safety Device Over temperature alarm Inner Chamber size(W*L*H)(mm)） 325*300*150 Exterior size (W*L*H)(mm) 350*318*210 Packing size(W*L*H)(mm) 420*390*280 Volume 14.6L Load per rack 5kg Shelf Number 1 Power Supply（50/60Hz）Current rating AC220V/4.5A NW/GW (kg） 6/7
81 | water distiller machine | pcs | 5 |  | 0 | Production Capacity: 20 L per hour (specify exact capacity) Input Water Requirements: Tap water or pre-filtered water, Temperature: 10–35 °C, Pressure: 0.1–0.3 MPa Distillation Method: Boiling of feed water → steam condensation → collection of distilled water, and continuous or batch mode Purity Specifications: Resistivity: ≥1 MΩ·cm at 25 °C, Total Dissolved Solids (TDS): <10 ppm, Bacteria: Significantly reduced (boiling process) and pH: Effectively neutral Material & Construction: Boiling chamber: Stainless steel 304/316, Condenser: Stainless steel or borosilicate glass coil and Collection tank: Stainless steel or glass-lined Temperature & Operating Conditions: Feed water temperature10–35 °C and Ambient operating temperature15–40 °C Electrical supply: 220–240 V AC, 50–60 Hz Safety Features: Overheat protection / thermal cut-off, Auto shut-off on low water level and Indicator lights for operation status Maintenance: Easy to clean boiling chamber and condenser, Scale and mineral deposits removable and Replaceable filters (if pre-filter included) Noise Level: <50 dBA (typical during operation) Standard Supply: Water distiller machine (>20 L/hr), Power cord and plug, User manual and Optional pre-filter cartridge
82 | Semi-Automated Rotary Microtome | pcs | 1 |  | 0 | Semi-automated rotary microtome that produces consistent high-quality sectioning. Versatile, user-friendly 7” touchscreen, full-color TFT 7” display with PCAD (Projective Capacitive Technology) touchscreen. Button Memory function (MEM) for the retrieval of previous cutting settings. Equipped with a specimen orientation head that works with different specimen clamp types and an easy-to-operate change system with locking mechanism. Cutting-edge design of the specimen clamp that prevents the building up of dirt and facilitates operations with the microtome. The knife carrier can be moved sideways both to the right and the left for optimal use of the cutting edge. The standard knife carrier can be used for both low profile and high profile disposable blades Smooth-working handwheel to minimize muscle strain. The handwheel can be locked in any angular position. Big capacity and easy to mount waste tray. Special antistatic coating preventing paraffin adhesion and minimizing cleaning time. Ergonomic design that includes a surface for armrests. Section thickness Total range in (μm) 0,5 – 100 Thickness settings in (μm) 0,5 – 5 in increments of 0,5 5 – 20 in increments of 1 20 – 60 in increments of 5 60 – 100 in increments of 10 Trimming Total range in (μm) 1 – 600 Thickness settings in (μm) 1 – 10 increments of 1 5 – 20 in increments of 1 20 – 50 in increments of 5 50 – 100 in increments of 10 100 -600 in increments of 50 Blade Holder Type Two in One for high- and low-profile blades Max. specimen size 50x60x40 mm (w x h x d) Dimensions Width 465 mm; Depth with tray 540 mm; Height 298 mm Weight 33,9 kg
83 | Embedding Center | set | 1 |  | 0 | Features : The Embedding system is composed of three modules, main console and Cryo Console and Thermo Console. It is independent design, can put the position left-to-right or right-to-left, to meet different habits of customers. The large working surface allows user processing many samples at the same time. The temperature and work start time can be settable through the control panel. Main Console: 1．It has a unique embedding upper lighting system design, using LED luminous tube lighting with enough brightness to determine the embedding position of small tissues 2．Large-capacity wax melting working cylinder, which can meet 400 wax blocks for embedding at one time 3．The machine adopts fully automatic program control, and can preset 7 days of automatic start time. 4．This machine has an independent five-temperature control system. The temperature of the wax tank, wax nozzle, workbench, holding tank and tweezers block can be adjusted within the range of 0-99℃. 5．With manual operation sensitive micro switch and foot control switch for the operator to choose, it is the first domestic similar product. Cryo Console: 1. It is composed of a compressor cooled work surface held to chill the paraffin blocks rapidly. 2. The Teflon coated surface allows a minimum ice formation and easy cleaning. 3. Temperature is adjustable from ambient to - 20 centigrade. 4. It can be used as a stand alone unit near the microtome station. Thermo Console: 1. It is removable in order to facilitate cleaning operations. 2. Stand alone unit near microtome/embedding unit. 3. It may be placed randomly up to user operating habit. 4.Thermal Console holds processed tissue specimens, in any type of cassettes, awaiting paraffin embedment. Stores and warms specimens in a large paraffin tank.Working Surface. Technical Data : 1. Nominal supply voltage: 220V/50Hz 2. Maximum power draw: 800-1000W 3. Working temperature: from ambient to 99℃ increments 1℃ 4. Capacity: Paraffin reservoir: 6.0 Lt Cassette/mold tray: 1.6 Lt 6．Dimensions: Width: 600mm Depth: 650mm Height: 460mm 7．Weight: 25Kg Cryo Console specification: 1. Nominal supply voltage: 220V/50Hz 2. Maximum power draw: 180 W 3. Working temperature: from 0℃ to -20℃increments 1℃ 4. Dimensions: Width: 410mm Depth: 660mm Height: 450mm 5. Weight: 25Kg Thermal Console specification 1. Nominal supply voltage: 220V/50Hz 2. Maximum power draw: 500W 3. Working temperature: from 30-90℃ 4. Capacity : Paraffin tray: 5Lt 5. Dimensions: Width: 410mm Depth: 620mm Height: 440mm 6. Weight: 14Kg Accessories: Standard Setting List 1. Foot Switch 1 pc 2. Magnifier 1 pc 3. Fuse (5 A) 2 pcs 4. Power line 3 line Optional Accessories: 1. Temperature sensor 2. LCD Screen
84 | Stretching table | pcs | 1 |  | 0 | Highlights • Histology slide warmer • Staged slide warming area for 30 glass microslides • Integrated overheat security and safety thermostat • Digital temperature display • Temperature selection up to 80 °C • Programmable working hours Specifications • Temperature range : up to 80 °C for slide warmer • Slide warmer capacity: 30 microslides • Slide drying area o Aluminum, Capacity: 30 slides up to 80 °C • Control : microprocessor, programming of temperature, working day, working time, time and date • Weight (w/o accessories): 15 kg • Electrical parameters: 230 V; 50/60 Hz, Pmax: 340 W Art.No.: 13000510 , 115 V; 50/60 Hz, Pmax: 340 W Art.No.: 13000511 • Pmax: 340 W Standard Accessories • Basic instrument • Histology slide warmer (staged, aluminum) • Main Cable • Operation manual
85 | Tissue Floatation Bath | Pcs | 1 |  | 0 | Highlights • Histology water bath • Separable glass water bath • Integrated water bath Illumination (LED, white/blue) • Integrated overheat security and safety thermostat • Digital temperature display • Temperature selection up to 80 °C • Programmable working hours SPECIFICATIONS • Temperature range : up to 80 °C for water bath • Water bath capacity : 2.0 litres • Water bath illumination : LED (white/blue) • Control :microprocessor, programming of temperature, working day, working time, time and date • Dimensions [W x D x H] : 350 x 370 x 230 mm • Electrical parameters o Art.No.: 13000520  230 V; 50/60 Hz, Pmax: 340 W o Art.No.: 13000521  115 V; 50/60 Hz, Pmax: 340 W • Pmax: 340 W Standard Accessories • Basic instrument • Histology water bath (glass, LED illumination) • Operation manual • The device will include its own dedicated UPS or Stabilizer • Mains cable
86 | Laminar flow hood | pcs | 1 |  | 0 | External Dimensions:(W*D*H)(mm) 1930*710*1850,Internal Dimensions:(W*D*H)(mm) 1830*615*625,Control System Microprocessor controller with LCD display, graphical and digital display of the operating status of each function.Work Surface Height 750mm,Airflow Velocity Adjustable, 0.2~0.5 m/s,HEPA Filter Efficiency > 99.995% at particle size of 0.3 μm, Dimensions(W*D*H)(mm) 895*470*70, Quantity 2pcs,Pre-Filter Material made of washable polyester fiber, Dimensions(W*D*H)(mm) 1000*425*2 , Sound Emission ≤62dB(A),Fluorescent Lamp 6.5W*2, LED Type, Intensity ≥900Lx, UV Lamp 20W*2, With timing function, Interlock with front sliding sash. Emission of 253.7 nanometers for best sterilization effect.Power Consumption 300W, Doesn’t include external device connected via power socket, Cabinet Construction Main Body: Electro-galvanized steel with antimicrobial Powder Coating, Work Zone: Stainless steel, grade 304, Manual Sliding Sash: 5 mm UV Resistant Tempered Glass with counter-balance system and ergonomic handle, Side Windows: 5 mm UV Resistant tempered glass Electricity AC220V ±10%, 50/60Hz,1P, Standard Accessories Base stand*1;LED lamp*1(only 2 for LCB-V6F);V lamp*1(only 2 for LCB-V6F);Waterproof socket*2 (only 1 for LCB-V3F);Foma casters*4, Optional Accessories Wind speed sensor; Gas tap;Watertap;Casters with brakes;Pipette stand ;Hook and hook bar Weight（kg） 330
87 | Co2 Gas Cylinder | pcs | 1 |  | 0 | Purity: Laboratory/research grade CO2CO sub 2 CO2 typically has a purity of 99.8% to 99.99% or higher. Physical State: The CO2CO sub 2 CO2 is a liquefied compressed gas inside the cylinder. Working Pressure: The standard pressure at room temperature (around 21∘C21 raised to the composed with power C 21∘C or 70∘F70 raised to the composed with power F 70∘F) is typically about 57-65 bar (around 830-940 psi), though the maximum working pressure can range from 150 to 200 bar depending on the cylinder's design and standard. Material: Cylinders are commonly made from seamless steel or aluminum alloy. Capacity: Common laboratory sizes vary widely, often measured by weight of gas (e.g., 5 lb, 10 lb, 20 lb, 50 lb) or by water capacity in liters (e.g., 10L, 20L, 40L, and 50L). Dimensions & Weight: Dimensions (height and diameter) and total weight vary significantly with capacity. For example, a 9 kg (≈20is approximately equal to 20 ≈20 lb) CO2CO sub 2 CO2 cylinder might be around 1300mm1300 mm 1300mm high and 200mm200 mm 200mm in diameter with a total weight of around 50 kg (tare weight). Color Code: The cylinder body is often painted a specific color (e.g., grey) according to regional or application-specific standards. Valve Connection: The outlet connection type depends on the country's standards (e.g., BS 341 No. 8, DIN 477, CGA 320 in the U.S., etc.). Safety Standards: Cylinders must meet specific design and testing standards, such as DOT, TC, ISO, or EN regulations, and undergo periodic hydrostatic testing
88 | Electrical Loop Strelizer (for Biosafety Cabinet) | pcs | 1 |  | 0 | 220V, Sterilization takes only 5 to 7 seconds at optimum sterilizing temperature of 1,600°,Sterilizes loops and needles safely and conveniently
89 | Automatic AST reader | pcs | 2 |  | 0 | Image Acquisition: high-resolution 5MP advanced camera) to capture images of the entire Petri plate. Zone Diameter Measurement: precise, automated measurement of the inhibition zone diameters in millimeters (mm). using image analysis algorithms (e.g., radial profile analysis algorithm) that detect the edges of the inhibition zones and the antibiotic disks. Inhibition Zone Detection Accuracy: should demonstrate high accuracy in zone detection, as aligning closely with manual technician measurements >90% accuracy. Antibiotic Disc Identification: Automated identification of the antibiotic discs, often using printed codes or patterns, to ensure correct interpretation against the associated antibiotic and its concentration. Expert System Software: The reader must integrate with or include expert system software that interprets the measured zone diameters. Clinical Breakpoint Adherence: The software should apply current clinical breakpoints and interpret results (Susceptible, Intermediate, Resistant - S/I/R) according to established international standards such as those from the EUCAST website or the Clinical and Laboratory Standards Institute (CLSI). Data Management and Traceability: The ability to store results, track samples, and integrate with a laboratory information system (LIS) for seamless data flow and reduced manual entry errors. Plate Compatibility: Designed to accommodate standard agar plate sizes, typically 90mm or 150mm round plates. Physical Specifications: Compact, robust, and easy-to-clean benchtop device with a small footprint 19 x 19 x 21 cm for one model. Connectivity: Connection to a PC, often via a standard USB cable, for data processing and software interface.
90 | Control poll | pcs | 2 |  | 0 | Material: aircraft-grade aluminum or heavy-duty stainless steel to withstand the force of a large, aggressive dog while remaining manageable for the handler. Length: 4 to 6 feet (1.2 to 1.8 meters) for general use, providing a safe distance between the handler and the dog Cable/Noose: The noose should be made of a strong, plastic-coated wire cable (e.g., tested to 350kg/750lb loading) to prevent injury to the animal's skin and hair. The noose diameter should be adjustable to safely fit different sized animals, with a maximum circumference of around 11 inches (28 cm) in diameter. Locking Mechanism: With an automatic locking device Release Mechanism: With a quick or instant release mechanism (button or spring-loaded latch),. Grip: With an ergonomic, non-slip rubber handle to ensure a secure grip. Protective Features: With plastic sleeve or ""tooth guard"" over the lower part of the pole to protect the handler's hands or prevent the dog from damaging the metal if it bites the pole. (optional) "
91 | Catch net | pcs | 2 |  | 0 | Netting: Should be made from common materials like thick, treated nylon or polyester mesh, sometimes double-ply for added strength. Frame/Handle: Aluminum or stainless steel .with the metal parts coated for rust and corrosion resistance. Net Opening: 20-24 inches (approx. 50-60 cm) in diameter wide enough to easily enclose the dog. Net Depth: (around 47 inches or 1.2-1.5 meters for a large dog net) Handle Length: 40 inches to over 80 inches (approx. 1m to 2m+) Closure Mechanism: With a sliding sleeve on the pole Grip: The handle should have an ergonomic, non-slip grip for secure handling and control. Maintenance: The net material should be easy to clean and resist wear, UV rays, and weather conditions
92 | Automated microbiological analyzer (Vitek) | pcs | 1 |  | 0 | Function: Automated microbial identification and antibiotic susceptibility testing (ID/AST). Technology: Advanced colorimetry and turbidimetry Time to Results: ID/AST results in as little as 6 to 8 hours for most bacteria. Workflow: Minimal hands-on time, "load and go" system once the inoculum is prepared and linked via barcode in a cassette. Traceability: Fully traceable workflow with barcoded cards and samples. Database: Extensive, regularly updated database of over 15,000 bug/drug combinations, incorporating CLSI/EUCAST standards Capacities:Available in various card capacities (e.g., 15, 30, 60, or 120 for VITEK 2 XL) to suit different lab sizes Physical: Weight ranges from 75 kg (Compact 15/30/60) to 145 kg (VITEK 2 XL).
93 | Automatic pipetter (pipett gun) | pcs | 1 |  | 0 | LCD Display shows battery charge and coarse speed setting • Lightweight ergonomic design • Comfortable concave triggers with fine speed control • 8 hour continuous use per charge • For use with 1-100ml pipets • Can fill a 25ml pipette in 6 seconds • Touch sensitive triggers for fine speed control • Gravity dispense, with blow-out for TD pipettes • Comes with 2 x 0.45um hydrophobic filters • Autoclavable nose-cone • Powerful lithium-ion battery, no 'memory' • Can be used while charging, no down-time • The LCD display shows the battery level and coarse speed setting. • Touch sensitive comfortable triggers control the fine speed of aspiration and pipetting process. • A replaceable 0.45um hydrophobic filter prevents liquids from contaminating the inside of the unit.
94 | PCR Machine (with accessories) | pcs | 3 |  | 0 | Features: Should have Programmable thermal gradient: for optimization of annealing tempratures; Large touch screen: for streamlined protocol programing and editing; Heated lid: for use with various vessels and sealer types; Expanded connectivity: WiFi, Ethernet and USB; Automated Lid: For integration with automation robotics; Network storage drive access for excellent data management; Application and Uses: ideal for Amplification/PCR; Cloning; Cycle Sequencing; Gene expression studies; Mutagenesis; Sample block type: Fixed; Memory: Not less than 4.6 GB; Noise Level: should be less than 90dB; Lid type: Automated; Lid heating: 30-110°C; Temprature range: 4-100°C; Thermal accuracy: +/- 0.2 °C of programmed target at 90 °C; Thermal Uniformity: +/- 0.4 °C well-to-well within 10 sec of arrival at 90°C; Gradient operational range: 30-100°C; Gradient programmable span: 8 rows, 1-24 °C; Touch screen display: 8 inch LCD; Programming options: Step-based graphical and automatic; Security features for regulated envrionments: Password protection and secure mode; Reporting: Exportable run logs and system logs; Temprature control modes: Calculated and block; Communiations: should have at least WiFi, USB 2.0, Ethernet; Heating and cooling method: Peltier; Lid Power Supply: 220-240 V, 50/60 Hz with Type F power plug; Electrical Approval: Should be supplied with IEC, CE Compliane Certificate. Accessories: CFX OPUS Ethernet Cable Real-Time PCR System to ethernet ports; Tube Frame: to provide extra support for individual 0.2ml tubes in the dual 48 and 96 well fast reaction modules; Tube Frame: to provide extra support for individual 0.2 ml tubes in the 96 deep well reaction module; PTC Tempo Thermal Cycler WiFi Adapter: WiFi antenna enables WiFi connection for PTC Tempo thermal cycler worldwide and should be supplied with recommended accessories that can operate the machine at least for 2 years; Warranty Period: not less 2 years

## Sheet2 (hidden)
Completed | Laboratory consumables | CAPEX | Board of Trustees | Local
Not Stated | Office Supplies | OPEX | Director General´s office | Foreign
On progress | IT Equipment and Services |  | Internal Audit
| Furniture and Fixtures |  | Institutional Planning Directorate
| Construction and Building Materials |  | Director Corporate services
| Professional Services |  | Finance Nairobi
| Vehicles and Transportation Services |  | General costs Nairobi
| Printing and Publishing |  | Resource mobilization unit
| Catering and Food Services |  | Human Resources Directorate
| Others |  | Liaison & Protocol Office
Overhead Recovery
DG’s Representative- Ethiopia
IP/ Legal management
Hostel-nbo
Hostel-Addis
Residential Housing NBO
Catering-add
Transport-Nbo
Transport-Addis
ILRI club - Addis
External Income-NBO
Investment Income
External Income-Addis
Residential Housing ADDIS
Other Center Income
Gym & Recreation
Unrestricted Funds
Catering and Conferencing Nbo
Capex Lab
Global ICT management
Computing services - Addis
OCS Training
Staff Welfare and Development
Human Resources-Nairobi
Human Resources-Addis
Platforms
Innovations Profiling and Scaling (IPS)
Research support- Nigeria
Research support- India
Research support- Vietnam
Research Support - Burkina Faso
Research support-Philippines
Research support-Zimbabwe
Research Support - Tanzania
Research Support -Nepal
Research Support -Senegal
Research support- Uganda
Health of Animals, People and Ecosystems (HEALTH)
Livestock Genetics, Nutrition and Feed Resources (LGN&FR)
Livestock, Climate, and the Environment (LCE)
People Policies and Institutions (PPI)
RS Recovery
DDG-Research Intergrated
DDG-Research Bios
Data and Research Methods Unit
Environment Health and Safety Safety
Capacity Development - Nairobi
Communications & Knowledge Management
DDG (Research) Office Biosciences
Animal Unit
Biorepository
Tick Unit
Vaccine Stabilate Diluent
Platform - SeGoLiP Group
Platform - Lab manager
Genomics Platform
BecA-ILRI Hub
Genebank Analytical Services
Genebank Analytical Services HSU
Mazingira Lab
Engineering Nbo
Engineering Addis
Security-nbo
Security-Addis
Stores-nbo
Stores-Addis
Procurement-Nbo
Procurement-addis
Liaison &Travel Addis
Office services - Add
Operations Nairobi
Finance Addis
Services Addis
Space Recovery-NBO
Space Recovery-Addis
Hosted Institute